## Slide 1
Automation
The simple meaning of automation is mechanization of various duties performed by a lab, personnel. These duties are :
Instructing patients
Collecting, labeling, separating and preserving various specimens such as blood, plasma serum etc.
organizing glasswares, reagents, instruments for specimen testing.
Pipetting reagents
Pipetting specimen, standard, caliberators or quality control sera.
Mixing reagents and specimen, caliberators or quality control sera respectively.
Incubating run minutes at specific temp.
Reading intensities of end products.
Printing the test reports.
Preparing quality control data and charts.

## Slide 2
Various advantages of Auto Analyzer
Large no. of samples can be tested in the short time.
Variety of tests can be performed by using various methods such as end point and kinetic.
By using techniques such as ELISA, Turbidimetry, chemiluminescene, nephelometry, it is possible to determine concentation of hormones, drugs, tumor markers and various types of abs. and proteins.
Most of the methods performed on automation are accurate, precises, sensitive and specific.
Although the various types of Auto Analyzer are expensive but in the long run they prove to be cast effective because the amount of reagent and specimens required can be as low as 300 and 5 ml respectively.
Automation allows laboratories to process much larger work load without a comparable inc. in no. of staff members.
Internet and external quality control programs can be implanted efficiency.
In the case of fully automated analyzers the lab staff members do not come in contact with specimens and reagents or biohazards materials and hence working on the autoanalyzers is very safe.

## Slide 3
History of Auto Analyzer
I successful automated system was introduced by L.T. Skeggs in 1957 and Technician and marketed it as single channel continuous flow analyzer.
Continuous flow Analyzers
Early automated chemistry analyzers were single or dual channel devices. In these system the samples and regent were passed sequentially through the same analytical pathway and separated y means of air bubbles. The relative proportions of samples and reagents were determined by their individual flow rates. Mixing occurred when tubes jointed to form a common pathway. The different parts of the system and their functions were as follows :
Probe : Its function was like a push button pipette.
Samples : This module was used to hold the batch of samples aviating analysis in separate cups on a circular tray which rotated at intervals.
Proportioning pump : This module determined the relative flow rates of samples and all the reagents.

## Slide 4
Dialyzer : It contained a semi-permeable membrane and when samples were passed through it, batches of protein free filtrate were obtained.
Heating Bath : This module was used to maintain reaction mixture at a constant temp. The batches of filtrates and reagents reacted to form coloured complexes at specific temp.
Colorimeter : This module contained a colorimeter with a flow through cuvette attached with a debubbler. The function of this module was the same as any photometer in a lab.
Printer: This module performed the function of recording the photometric readings and calculating the values of the analytic in a printing form.
Advantages of single channel continuous flow analyzer
It was possible to test large no. of specimen for a particular test like glucose or urea accurately precisely in a short duration.
Disadvantages:
At a time only on type of determination was performed by single channel continuous flow analyzer.
This auto analyzer occupied a larger spaces in the lab.

## Slide 5
Multichannel continuous flow analyzer
Techniqon introduced following Auto Analyzer subsequently which can perform 6-12 diff. tests simultaneously.
Sequential multiple analyzer (SMA 12/60)
This auto analyzer was used to process 60 specimens/hr. and reported 12 tests simultaneously. These tests were glucose, urea, creatinine, total proteins, total bilirubin, SGPT, SGOT, aDH, Alkaline phosphatase, uric acid, calcium, phosphorus.
Sequential multiple analyzer (SMA 6/60)
This auto analyzer perform 60 specimens/hr. and reported the results of 6 tests simultaneously. These tests were chlorides, CO2++, K, Na, urea, Glucose.
Disadvantages of SMA
Only 6-12 fixed determinations were performed by these analyzers and physicians might not be interested in all these tests and they might be interested in some other tests which the analyzer did not perform.
This auto analyzer occupied large space in the lab.
Advantages of SMA
The techniqon SMA system was designed to perform diagnostic assays of human serum for mx. 18 biochemical parameters simultaneously at the processing rate of 150 samples/hr. The system consisted functionally of an analytical processor and or result processor. The analytical processor performed and monitored the assay of each sample. The results processor provided extensive data. Data handling for sample results and patient informational.
It had the capacity of storing over 2000 sample records in the files.

## Slide 6
Discrete Auto Analyzer
Latest clinical chemistry analyzers coordinate multiple operations to a smoothly functioning system. This requires components such as specimen handling, reagent system, optical and computers to be fully integrated with one another.
Perhaps the most significant technology lead has been in the area of computers without these advances the full potential of newly developed specimen, reagent, liquid and optical handling system could not be realized.
The basis of computers is a miniature transistor on a silicon chip called as Microchip. The Microprocessor is the central processing unit of the digital computer and is contained on a single chip. The microprocessor perform multiple function on an instrument. They are responsible for data management operations such as optical reading, matching cup positions with results, collating patient results with data and reporting and retrieving data.
Microprocessor controls the precise timing of a robotic arms so that pipetting and dispencing corresponds to run, turn table movements. This ensures proper liquid handling and optical readings that match a given cuvette. Many instrument use several microprocessors each dedicated to perform specific functions to meet the needs of patients, physician and various complexity level lab. Such as low, moderate or high, discrete types of Auto Anayzers were introduced and soon they become extremely popular because they suited to all above mentioned categories of the lab.

## Slide 7
BASIC FEATURES
Reagent and specimen are pipetted in a cuvette mixed and the end product is recorded photometrically. For every specimen & reagent a separate cuvette is used.
All end point runs, two pt. assays and rate of run methods can be performed by these analyzers with monochromatic or bichormatic readings.
More than 100 opened test programs can be selected through the reybond.
These system are extremely economical since reagent quantity required is only 200-250 ml per test.
These analyzer after patient wise report print out.
Levy-jenining control programs can be performed.
Minimum 200 test results could be stored.
Apart from routine chemistry test is possible to perform assay to determine levels of hormones, drugs, electrolytes.
Even the simpler types of discrete auto analyzer offer user friendly software.