## Slide 1: High Performance Liquid Chromatography (HPLC)
What is HPLC?
Types of Separations
Columns and Stationary Phases
Mobile Phases and Their Role in Separations
Injection in HPLC
Detection in HPLC
Variations on Traditional HPLC
  Ion Chromatography
  Size Exclusion Chromatography

## Slide 2: Isocratic versus Gradient Elution
Isocratic elution has a constant mobile phase composition
  Can often use one pump!
  Mix solvents together ahead of time!
  Simpler, no mixing chamber required
  Limited flexibility, not used much in research
    mostly process chemistry or routine analysis.
    Gradient elution has a varying mobile phase composition
  Uses multiple pumps whose output is mixed together
    often 2-4 pumps (binary to quarternary systems)
  Changing mobile phase components changes the polarity index
    can be used to subsequently elute compounds that were previously (intentionally) “stuck” on the column
    Some additional wear on the stationary phase
    Column has to re-equilibrate to original conditions after each run (takes additional time).

## Slide 3: Mobile phase delivery system
With varied flow rates & pressures solvent is introduced in column
If system is not pulse free, pulse damping system, degasser is used to remove air/ gasses from solvent (mobile phase)
Capable of delivering solvent gradient
Pump operates at 100 atm 1500 psi to 400 atm to 6000 psi
0.5-2 ml/min to few µl/min (micro bore)

## Slide 4
There are a number of different types of pumps that can provide the necessary pressures and flow-rates required by the modern liquid chromatograph.
In the early years of the LC renaissance, there were two types of pump in common use; they were the
Pneumatic pump, where the necessary high pressures were achieved by pneumatic amplification (under pressure)
syringe pump, which was simply a large, strongly constructed syringe with a plunger that was driven by a motor.
Today the majority of modern chromatographs are fitted with reciprocating pumps fitted with either pistons or diaphragms.
Pumps

## Slide 5
The Pneumatic Pump
The pneumatic pump has a much larger flow capacity than the piston type pumps but, nowadays, is largely used for column packing and not for general analysis. The pneumatic pump can provide extremely high pressures and is relatively inexpensive, but the high pressure models are a little cumbersome and, at high flow rates, can consume considerable quantities of compressed air.
Air enters port (1) and applies a pressure to the upper piston that is directly transmitted to the lower piston. If connected to a liquid chromatograph, or column packing manifold, liquid will flow out of the left hand side non-return valve as shown in the diagram which continues until the maximum movement of the piston is reached. At the extreme of movement a micro switch is activated and the air pressure transferred from port (1) to port (2).
The piston now moves in the opposite direction and draws mobile phase through the right hand non-return valve refilling the cylinder with solvent. Again on reaching the limit of the piston movement a second micro switches the air supply from port (2) back to port (1) and the process repeated. The refilling process is extremely fast and, if an appropriate pulse dampener is used, the outlet pressure remains reasonably constant during the refill cycle.

## Slide 6
Non-Return Valves
For efficient function, it is important that while the piston compresses the solvent to express it from the exit port, the flow is completely stopped at the inlet port.
Conversely, when the pump draws fresh solvent into the cylinder during refill, the non-return valves must allow solvent to flow through the inlet valve but, flow-back from the exit valve must be completely stopped. This is achieved by the use of efficient non-return valves.
The critical part of the valve consists of a synthetic sapphire ball resting on a seat. The seat may be of stainless steel, captured PTFE or, more commonly, also of sapphire. When the flow is directed against the ball the ball moves forward allowing the liquid to flow pass it. When the direction of pressure changes resulting in potential flow-back through the valve, the ball is driven back onto its seat and stops the flow.

## Slide 7
The Single Piston Reciprocating Pump
The single piston reciprocating pump was the first of its type to be used with high efficiency LC columns (columns packed with small particles) and is still very popular today. It is simple in design and relatively inexpensive.
During this movement the cylinder is loaded with more solvent through the inlet non-return valve. The shape of the cam is cut to provide a linear movement of the piston during expression of the solvent but a sudden return movement on the refill stroke. In this way the pulse effect that results from the refill action is reduced. The pulses, however, are not completely eliminated and the detector noise resulting from these pulses is probably the most serious disadvantage of the single piston pump.
Most pistons of modern LC pumps are made of synthetic sapphire to reduce wear and extend the working life of the pump. The cylinder is usually made of stainless steel and is attached to two non-return valves in line with the inlet and outlet connections to the pump. The piston is driven by a stainless steel cam which forces the piston into the cylinder expressing the solvent through the exit non-return valve. After reaching the maximum movement, the piston follows the cam and returns as a result of the pressure exerted by the return spring.

## Slide 9
Twin Piston Pump
The twin piston pumps with short
stroke are among the most commonly
used pumps for HPLC.
Both pump heads are switched in
series, whereby the piston in the first
pump head delivers a specific volume
per stroke. One pump is being filled while the other is delivering the solvent.
An excentric disk presses piston 1 to the right and displaces the solvent. The double ball saphire valves ensure that the solvent stream can flow in only one direction.
The second piston is used to produce a nearly complete pulsation damping. With the twin piston pump, a pressure of 40 MPa is achieved.

## Slide 11: Reciprocating Piston pumps
Inexpensive, wide range of flow rates
Small motor driven piston, moves rapidly back n forth in hydraulic chamber.
Capacity 35-400µl.
Check valves on background stroke piston sucks solvent from MP reservoir. Outlet to separating column is closed.
On forward stroke, pump pushes solvent to column. Inlet from reservoir is closed.
Hydraulic fluid allows pumping action to solvent via diaphragm. (minimizes contamination & corrosion of pump parts)

## Slide 12: Reciprocating Pump
Stroke vol during each cycle/stroke frequently varied for flow rates.
Dual head/triple head pumps – identical piston chamber, but solvent delivery is smoothed bcos one piston sucks & other delivers the solvent.
Solvent delivery is maintained continuous.
No restriction on solvent reservoir size/operating time.

## Slide 13: Syringe type pumps
Work on positive solvent displacement by piston mechanically driven at const rate.
Piston is actuated by screw feed drive in geared box run by motor.
Rate of flow can be changed by changing voltage of motor.
Finite capacity solvent chamber 250-500ml
Good for small bore columns.
Pulseless flow possible with pressure capacity 200-475 atm.
Constant pressure pump

## Slide 14
The Syringe Pump
The syringe pump is a large, electrically operated simulation of a hypodermic syringe. Although used in the early days of LC renaissance, it is rarely used today as, due to its design, it can provide only a limited pressure and the volume of mobile phase available for use is restricted to the pump volume. Unless the separation is stopped while the pump is refilled and the development subsequently continued, the pump can only elute solutes that have retention volumes equal or less than the pump capacity.

## Slide 16: Pulse dampers
Many detectors r sensitive to variation in flow.- R.I., Electrochemical, conductivity, fluorescent, U.V. absorbance detectors.
Exhibit ↑ detector noise with flow modulation.
Simple damping method: flexible bellows/compressible gas/fluid in capped up right portion of tee tube & takes some of pulsation energy.
When pump refills this energy is released to smooth the pressure pulsation.
Have large fluid volume.
Excess pressure of 100psi r req for effective operation.
Electronic pulse dampers provide small rapid forward stroke of piston & rapid refill stroke.
Dampens the pulse flow.
Solvent is brought upto system pressure before delivery.

## Slide 17: Sample introduction
Usually 5 to 1000 L volumes, all directly onto the column
  not much worry about capacity since the columns have a large volume (packed).
Injector is the last component before the column(s)
A source of poor precision in HPLC
  errors of 2-3 % are due to injection
  other errors are added to this
  due to capillary action and the small dimensions/cavities inside the injector
6-PORT Rotary Valve is the standard manual injector
Automatic injectors are available
Two positions, load and inject in the typical injector
Injection loop internal volume determines injection volume.

## Slide 18: HPLC - Injector
Six-port Rheodyne valve in which the sample fills an external loop
A clockwise rotation of the valve rotor places the sample-filled loop into the mobile-phase stream, with subsequent injection of the sample onto the top of the column

## Slide 19
Inject (move the sample
loop into the mobile
phase flow)
LOAD (the sample loop)
http://www.restek.com/Technical-Resources/Injector-Animations

## Slide 20
Micro sampling injector valves:
sample dissolved in MP to avoid solvent peak.
Calibrated sample loop is filled by flushing it thoroughly with sample solution by means of ordinary syringe.
Rotation of the valve rotor places filled loop into MP stream with subsequent injection of sample onto the top of the column.
Sample loops of different cap- 10- 20 µl.