## Slide 1
Discrete Auto Analyzer
Latest clinical chemistry analyzers coordinate multiple operations to a smoothly functioning system. This requires components such as specimen handling, reagent system, optical and computers to be fully integrated with one another.
Perhaps the most significant technology lead has been in the area of computers without these advances the full potential of newly developed specimen, reagent, liquid and optical handling system could not be realized.
The basis of computers is a miniature transistor on a silicon chip called as Microchip. The Microprocessor is the central processing unit of the digital computer and is contained on a single chip. The microprocessor perform multiple function on an instrument. They are responsible for data management operations such as optical reading, matching cup positions with results, collating patient results with data and reporting and retrieving data.
Microprocessor controls the precise timing of a robotic arms so that pipetting and dispensing corresponds to run, turn table movements. This ensures proper liquid handling and optical readings that match a given cuvette. Many instrument use several microprocessors each dedicated to perform specific functions to meet the needs of patients, physician and various complexity level lab. Such as low, moderate or high, discrete types of Auto Anayzers were introduced and soon they become extremely popular because they suited to all above mentioned categories of the lab.

## Slide 2
BASIC FEATURES
Reagent and specimen are pipetted in a cuvette mixed and the end product is recorded photometrically. For every specimen & reagent a separate cuvette is used.
All end point runs, two pt. assays and rate of run methods can be performed by these analyzers with monochromatic or Bichromatic readings.
More than 100 opened test programs can be selected through the Keyboard.
These system are extremely economical since reagent quantity required is only 200-250 ml per test.
These analyzer offer patient-wise report print out.
Levey-jening control programs can be performed.
Minimum 200 test results could be stored.
Apart from routine chemistry test is possible to perform assay to determine levels of hormones, drugs, electrolytes.
Even the simpler types of discrete auto analyzer offer user friendly software.
Various types of Discrete Auto Analyzer
The discrete type of auto analyzer are of 2 types :
Semi automated
Fully automated
Fully automated can further be divided into
Batch analyzer
Random Access Analyzer

## Slide 3
Semi Automated Discrete Analyzer
These analyzer are called semi auto analyzer because the initial stages of a specimen analysis are performed by the lab. Technician. These stages are pipetting of reagent, pipetting of specimen, mixing and incubating the reaction mixture. Following functions performed by the semi auto analyzer :-
Reading and point, rate of run, visible or U-V kinetic in monochromatic and bichromatic modes.
Displaying test results, printing and memorizing these results and graphs of all linear and non-linear runs.
Various specific features of Semi Auto Analyzer
Allows all routinely performed and all types of special in clinical biochemistry to program directly through keyboard.
Offers economical methodology where only 200-500ml of the reagent used per test.
User friendly software and offers backlit LCD display.
Lists all the methods stored in the memory.
Plots absorbance v/s wavelength for selected range.
Displays and prints system parameters.

## Slide 4
Performs statistical analysis and determinants mean values, standard deviation (SD), cumulative value for central samples and also plots the levey-jenning charts.
Performs immune assays, coagulation and drug assays.
Works without power by battery backup system.
The semi auto analyzers are extremely suitable for medium complexity lab.
Three different versions of semi-auto analyzer are :
1. Analyzer with built in dry block at 370C with about 12 cuvette positions and one for reading. These cuvettes are handled by manual operations.
Eg. STAT fx 2000, POINTA 180+ and MINILAB.
2. Analyzers with built in fixed metal thermo cuvette and operated by automated aspiration and washed by peristaltic pump are recorded by the analyzer.
Eg. METRO LAB 1600 plus, MICRO LAB, BIOTRON BTR 810.
3. Analyzers with following triple cuvette systems, 33 ml flow cell cuvette, manual square cuvette, manual round tube for coagulation and ELISA.
Eg. ERBA Chem 5 and ERBA Chem Pro.

## Slide 5
FULLY AUTOMATED DISCRETE ANALYZERS
These autoanalyzers perform all the functions of a semi-autoanalyzer and in addition to that they also perform following additional functions :
Automated dispensing of reagents.
Automated dispensing of samples.
Automatic mixing of reaction-mixtures.
Incubating the reaction mixtures.
These analyzers are supplied with the following additional features: -
A sampler to accommodate various samples standards, calibrators and quality control serum.
Programmable sample probes which identity sample pipettes sample vol. as 3ml with precision. The sample probes are provided with level sensors to detect the level of samples.

## Slide 6
Reagent probes for the pipetting of reagent. These are also equipped of regents sensors and dynamic reagent monitoring which keeps the operator continuously inform on residual reagent vol. at any pt. of time.
Reagent carousel (cuvette containers) this according various reagents container.
Peristaltic pump, water & waste bottles with level sensors for every excitability & maintain.
Reusable living cell in convenient centrifugal.
Programmable `wash cycles b/w samples & tests for min. carry over.
Cuvette rinsing station : The advanced versions of fully automated analyzers offer this facility to clean each cuvette with detergent and deionised water before use. This ensures excellent precision and reliability of results.
The fully automated analyzers can be of 2 types.
Batch analyzers
Random access Analyzer

## Slide 7
BATCH ANALYZERS
The auto analyzers of this type have become outdated. These analyzers perform only one type of test at a time. There was provision to accommodate only one type of test or only one type of reagent at a time. These analyzers were utilized by using following general procedure :
After collecting all time samples, particularly blood respective sera or plasma were separated.
Various batches of the samples were made according to the determinations such as glucose, urea, creatinine, protein etc.
Specific reagent was placed on the analyzer.
The analyzer was programmed according to the no. of specimens on the sampler and specific code no. was given for the test.
The analyzer performed following function :
Pipetting of reagent in specific cups or cuvettes.
Pipetting of sample in the reagent.
Automatic mixing of the reagent & sample.
Incubation of reaction mixture at a specific temperature.
Aspiration of the reaction mixture into the reading station.
Reading, calculating, displaying and printing of test results.
For eg. Clinicon. c

## Slide 8
Advantages
Large no. of samples batches were tested accurately and precisely with appropriate quality control program.
Disadvantages
These analyzers were not patient oriented and did not function according to the patient test.
These analyzers were not equipped with ‘STAT’ facility, immediate reporting or emergency testing of a patients sample.
These analyzers were not equipped with facility for Random Access Analysis.
Eg. ERBA-Chem10, Clinicon Corona and auto Pacer.

## Slide 9
RANDOM ACCESS ANALYZERS
These analyzers perform all the functions of batch analyzer, but in addition to that they are also equipped with :
Random Access Mode : Completing all test on one sample before proceeding to the next sample.
Sample orientation mode : Completing the greatest no. of smaples in the shortest period of time, that may be 120-800 samples per hour.
Sequential mode : Processing one test at a time.
Compared to batch analyzer, random access analyzers are equipped with following additional facilities:
Cuvette disc with temp. control and automatic multistage cleaning system.
Reagent table with capacity for single or double reagent containers and facility for low temp. 8-150C to preserves the stability of the reagents.
Level sensors for samples and reagents.

## Slide 10
Sample rack system, individual racks for samples, controls, calibrators.
Barcode identification of samples and reagents.
Facility for continuous loading of samples.
Facility for auto dilution.
Plotting of daily and monthly quality control charts.
Availability of ion selective electrodes, ISE Vol. for the determination of NA K & fluoride.
Capability to perform 2-3 reagents tests.
For eg. “ ERBA EXCEL 600, HITACHI 704, 705, ABBOT SPECTRUM, 550 EXPRESS; KODA DRY CHEMISTRY ANALYZER, ERBA SUPER STAT919, ERBA SMART LAB.
The component steps in fully automated analyzer system
Specimen Identification
A bar coded label is placed on the specimen container and it is read subsequently by one or more bar code readers that have been placed at key positions in the analytical pathway. The sample identification and information is then transferred to and processed by the system software. By using this method of specimen identification, the error rate is reduced to 1 in 300 characters or about in 10 in 1 million characters. By using bar code information, identification at patient’s bed side is possible to a greater integrity.

## Slide 11
Specimen preparation, handling and delivery
Adequate quantity of blood is collected in vaccutainer tubes and after proper centrifugation, the tubes are placed in the sample carousals. Since, the specimen probe is equipped with levels sensors, the required amount of serum can be pipette automatically.
The use of level sensors, restricts the penetration of sample probes into specimens, provides smooth motion and eliminates spatter. Disposable specimen cups are used, which are made up of glass or polyvinyl plastic. Their shapes are such that even without a cap, little evaporation occurs. The specimens are advanced into the analyzer at a uniform rate. The design of the device for holding specimen is often specific. It may be in the form of tray rack or chain. Carryover b/w specimens is reduced considerable by setting adequate wash stations.

## Slide 12
Reagent handling, storage identification & delivery
Liquid reagents stored in plastic or glass containers are used in most automated systems. Although, for most of the test single reagent monostep method is used, two or more reagent may be required for same tests. In large auto analyzers system such as HITACHI 747 reagents are stored in a separate compartment maintained at 4-100C. Reagents prepare manually or by automated reconstitution from dry tablet or powder are stable after reconstitution for 2-8 days when stored at 4-80C. Levels on reagent container include information such as reagent identification, volume of the contents expiry date and lot no. Many reagent contains carry bar cods that contain most of this information when a bar code reader is coupled with a level sensing system on the reagent, it alerts the operator about the quantity of the reagent in the container.
Liquid of reagents are taken up and delivered to mixing & run, cups & cuvettes by pumps via tubes or by positive displacement syringe devices. Syringe devices are commonly used in many automated systems for both reagent and sample delivery. Reagents probes are wash and pushed if more than one reagent is used to prevent reagent carryover.

## Slide 13
Mixing, Incubating & Reading
The mixing of specimen & reagent takes place in specific cells, which are also used to mix reactants like in thermocuvettes. Various techniques used to mix reactants include forceful dispensing magnetic stirring by using vibrating rods rotating pedal use of ultrasonic energy. For incubation the analysis. Compartment is placed in a large water bath or Air bath heated 30-370C. The cuvettes are transparent in the spectrum. The use of disposable cuvettes eliminates carryover. The cuvettes are washed in oiled & alkaline solution & then rinsed by deionised water & dried by vacuum.
Measurement devices :
For the measurement of results, principles are used which are based on the following techniques:
Photometry/spectrophotometery, reflective photometry.
Turbidimetry, Nephlometry
Ions selective electrodes or other electrode chemical techniques.

## Slide 14
For Photometric and spectrophotometeric reading 3 basic components are required
A radial energy source
Means of spectral isolation
Detector
These radiant energy sources used are tungsten, quartz, halogen, deuterium, mercury, neon and loser sources.
Spectral isolation is achieved with interference filters in several multi test analyzers, the required filters are mounts in a filter wheel & the appropriate filter is moved into the place under the command of microprocessor. Monochromators with moveable gratings & slits provide a continuous choice of wavelength.
Photo diode and photo multipliers tubes are used as detectors in most of the automated system. For each measured analyte, the microprocessor or chooses the appropriate pair of wavelengths. The readings are noted at two wavelengths. This eliminates the interference die to hemolysis, turbidity or high bilirubin concentration. Perp. Light path is maintained through the reaction cells or cuvettes by accurate alignment of lenses. The run cells and cuvettes are transported passed the light beams in a uniform manner in the case of reflectance photometry diffused reflected light is measured.

## Slide 15
Turbidimetry, Nephlometry
In the case of turbidimetry and nephlometry the light transmitted by the turbidity by the reaction between reagent and sample is measured.
These techniques are used for the determination of various protein and Ab in the sample. The composition of components incorporated into an automated analyzer is generally appropriate to the intended through put precision and range of wavelength and cast of the instrument.
Ion selective methodology is used by many automation analyzers for the determination of sodium, potassium, fluoride and lithium. Peristaltic pumps are used to move the sample into the chambers containing fixed and reference electrode. The specimen remain in contact with the electrode for sufficiently long time 7-45 sec. to reach steady state condition.

## Slide 16
Signal processing, handling of data and control of microprocessor :
The digital computers in automated system convert analogue signals from detectors to digital form by analog to digital converter the computer then processes the digital data by means of algorithms immediately in to useful output the microprocessor perform following various functions.
(A) All electromechanical operations are performed uniformly repeatedly and, in corrected sequence. They include transfer of solution, placement of proper filter by bichromatic and multichromatic reading, regulation of change of speed of rotation of the carousels addition to the function the microprocessor also control the function such as result calculations and result reporting by mean of digitals display and by printing.
(B) Microprocessor acquire, process, access and store operational data from the analyzer they also react to in proper function by recording the right nature of main function.

## Slide 17
(C) Important communication interactions related to trouble shooting are made possible by microprocessors. The signals may instruct the operator to intervene to replenish reagents supply to empty waste containers or reading operational problems.
(D) Output data tested against pre-set criteria can be flagged and displayed for the operators evaluation and judgment such information specifies that
a. linearity of a reactionis maintained a the run is non-linear
b. abs. of end product is very high.
c. abs. of end product is every high.
(E) The microprocessors can take command from the operator through the keyboards of touch sensitive screen. These common may be program temp. set pt., cal. Concentration timing intervals of kinetic run., format of printed output and quality control charts.