## Slide 1: UNIT V
Calibration & Validation
Introduction, definition & general principles of calibration, qualification & validation, importance & scope of validation, types of validation, validation master plan, calibration of pH meter, qualification of UV-visible spectrophotometer, General principles of analytical method validation
Warehousing: Good warehousing practice, materials management

## Slide 2: Calibration & Validation
Introduction:
Concept of Validation: Ted Byers & Bud Loftus, U.S. FDA officials, in mid 1970’s, more precise 1978 to improve the quality of pharmaceuticals
Drug regulatory authorities: ask for process, procedures, intermediate stages of inspections, testing, adapted during Mfg. are adapted to produce similar, reproducible, desired results meeting quality standards .The procedures are developed through validation to maintain or assure high degree of quality.
VALIDATION: essential element of Pharm. quality assurance.
Validated process provides documented evidence & high degree of assurance that uniform batches meeting the required specs. will be consistently produced.
Definition of validation: As per WHO: Action of proving & documenting that any process, procedure or method actually & consistently leads to expected results.
Validation is the process of evaluating products or analytical methods to ensure compliance with products or cleaning method requirements.

## Slide 3: Qualification
It is action of providing documented evidence that the equipment, premises, and systems are designed, installed and perform correctly & are ready for intended use. Qualification if a part of validation.
There are six stages of qualifications:
Design qualification (DQ): It provides documented verification that all key aspects of the design, procurement & installation adhere to the approved design intention & that all the manufacturer’s recommendations have been suitably considered.
b) Installation Qualification (IQ): Document that the equipment is properly installed according to the manufacturer and purchaser’s specifications. It covers equipment/system descriptions, which includes principle of operation, design requirements, equipment specifications piping, instruments diagrams, facility functional specifications, equipment utility requirements, and equipment safety features.

## Slide 4: Qualification
c) Operational qualification (OQ):
Document that the equipment operates within established limits and tolerances. It covers equipment operation procedures established and challenged equipment control functions, calibration requirements and schedules established, and maintenance requirements.
d) Performance qualification (PQ):
Documented evidence that the equipment can operate reliably as intended for the process under routine, minimum, and maximum operating ranges. To prove consistency in the performance of the equipments etc. test results are to be collected over a suitable time period.

## Slide 5: Qualification
e) Component qualification (CQ): It refers to manufacturing of auxiliary components to ensure that they are manufactured to the correct design criteria. CQ is performed for the components such as cartons, labels, shipping cases, phase change material etc. Random inspections are to be carried out at the third party’s Mfg. site to ensure that the components meeting the required specs are produced
f) Re- qualification: Refers to ensure that the equipments is still in the qualified state, after any modification of relocation of the equipment is done. Re qualification is also important at specific time intervals.

## Slide 6: Importance of Validation
Quality assurance: Validation ensures that quality is built into the system at every step. It checks accuracy & reliability of processes, procedures, methods to meet predetermined criteria. It provides high degree of assurance & maintain batch to batch quality
Cost reduction: Validation Helps in optimization of processes. It reduces sampling & testing procedures within routine production, hence less re-inspection, retesting, rework, rejection, & reduced no. of complaints & recalls. Thus save cost.
Regulatory compliance: It is an integral part of GMP guidelines. So, compliance is must.
It assures quality
Customer./patient satisfaction
Quality is inbuilt into product
Increased efficiency, shortening lead time resulting in low inventories
Fewer rejects,reworks

## Slide 7: Importance of Validation
9. Longer equipment life by operating as per specifications & establishing of cost effective preventive schedules
10. Government regulations, successful inspections, approved products
11. Ability to export, increased throughput, reduction in utility cost
12. Avoid capital expenditures, fewer complaints about process related failures.
13. Reduced testing in process and finished goods.
14. More rapid & accurate investigations into process deviations
15. More rapid & reliable startup of new equipment.
16. Easier scale up from development work
17. Easier maintenance of the equipment
18. Improved employee awareness of processes
19. More rapid automation

## Slide 8: PRINCIPLES OF VALIDATION
the execution of validation should be in compliance with regulatory expectations;
quality, safety and efficacy must be designed and built into the product;
quality cannot be inspected or tested into the product;
quality risk management principles should be applied in determining the need, scope and extent of validation;
ongoing review should take place to ensure that the validated state is maintained and opportunities for continuing improvement are identified.

## Slide 9: SCOPE of Validation
Pharmaceutical Validation is a vast area of work and it practically covers every aspect of pharmaceutical processing activities, hence defining the Scope of Validation becomes a really difficult task. However, a systematic look at the pharmaceutical operations will point out at least the following areas for pharmaceutical validation;
Analytical
Instrument Calibration
Process Utility services
Raw materials
Packaging materials
Equipment
Facilities
Manufacturing operations, Product Design, Cleaning. Operators

## Slide 10: SCOPE of Validation
Selection of raw material i.e. raw materials of desired quality attributes
Product design based on expected performance
Process design to build desired quality attributes in product
Design of control parameters like change control, acceptance criteria, tolerance limits etc.
In process quality control parameters & sampling plans
Finished product testing or evaluation criteria
Validation of related analytical process
Validation of related system, facility & equipment
Personnel training
Validation involves careful determination of criteria variable of the process. Such as moisture content of granules, drying temperature of time, etc. and then establishment

## Slide 11: PHASES OF VALIDATION
The activities relating to validation studies are classified into three phases.
Phase 1:
Pre-validation phase or the qualification phase ,which covers all activities relating to product R & D, formulation, pilot batch studies, scale-up studies, transfer of technology to commercial scale batches, establishing stability conditions, storage and handling of in-process and finished dosage form, equipment qualification ,installation qualification, master production documents, operational qualification, process capability.
Phase 2:
Process validation phase (process qualification phase) designed to verify that all established critical process parameters are valid & satisfactory products can be produced even under the worst case conditions.
Phase 3:
Validation maintenance phase requiring frequent review of all process related documents ,including validation audit report to assure that there have been no changes ,deviations, failures modification to production process, and that all SOP’s have been followed including change control procedures. At this stage the validation team also assures that there have been no changes /deviations that should have resulted in requalification and revalidation.

## Slide 12: TYPES OF VALIDATION
1. Process validation: is establishing documented evidence which provides a high degree of assurance that a specific process (such as the manufacture of pharmaceutical dosage forms) will consistently produce a product meeting its predetermined specifications and quality characteristics.
Stage of activities in Process validation :
Stage 1 – Process Design: The commercial manufacturing process is defined during this stage based on knowledge gained through development and scale-up activities.
Stage 2 – Process Qualification: During this stage, the process design is evaluated to determine if the process is capable of reproducible commercial manufacturing.
Stage 3 – Continued Process Verification: Ongoing assurance is gained during routine production that the process remains in a state of control.

## Slide 13: TYPES OF VALIDATION
Types of Process Validation:-
1. Prospective Validation: It is establishment of documented evidence of what a system does or what it purports to do based upon a plan. This validation is conducted prior to the distribution of new product.
2. Retrospective Validation: It is the establishment of documented evidence of what a system does or what it purports to do based upon the review and analysis of the existing information. This is conducted in a product already distributed based on accumulated data of production, testing and control.
3. Concurrent Validation: It is establishment of documented evidence of what a system does or what it purports to do information generated during implemented of the system.
4. Revalidation: Whenever there are changes in packaging, formulation, equipment or processes which could have impact on product effectiveness or product characteristics, there should be revalidation of the validated process.

## Slide 14: TYPES OF VALIDATION
Cleaning validation: is the process of assuring that cleaning procedures effectively remove the residue from manufacturing equipment / facilities below a predetermined level. Cleaning validation primarily applicable to the cleaning of process manufacturing equipment in the pharmaceutical industry. The term cleaning validation is to be used to describe the analytical investigation of a cleaning procedures or cycle. It should also explain the development of acceptance criteria, including chemical & microbial specifications, limits of detection & the selection of sampling methods.
The reasons for validating the cleaning procedure:
It is customer requirement. Ensures safety & purity of product. It is regularity requirement in API manufacture. Pharmaceutical products & API can be contaminated by other pharmaceutical products, cleaning agents & microbial contamination. The objective of the cleaning validation is to verify the effectiveness of cleaning procedure for removal of product residues, degradation products, preservatives, excipients and/ or cleaning agents as well as the control potential microbial contamination.

## Slide 15: TYPES OF VALIDATION
Elements of Cleaning Validation
Establishments of acceptance criteria : procedure consistently removes residues of substance previously manufactured down to levels acceptable and the cleaning procedure itself does not contribute unacceptable levels of residual materials to the equipment.
b) Cleaning procedure
c) Residues to be cleaned: Cleaning limits, Solubility's of the residues
d) Cleaning agent parameters to be evaluated: solubility properties, Environmental, Health and safety considerations
e) Cleaning techniques to be evaluated: Manual cleaning, CIP (Clean-in place), COP(clean-out-of-place), Semi automatic, Automatic, Time considerations
f) Other requirements: Sampling Techniques: Direct surface sampling:, Swab sampling, . Rinse sampling, Placebo Sampling Method

## Slide 16: TYPES OF VALIDATION
3. ANALYTICAL METHOD VALIDATION:
Refers to steps necessary to perform analytical test. Analytical methods needs to be validated and re validated as the conditions of test may change.
It is the process to establish the performance characteristics of analytical method fulfilling the required analytical applications.
It demonstrates that the method is suitable for intended purpose
Validity of specific method can be demonstrated in lab using samples/ standards similar to unknown samples.
Role of each departments involved in validation must be clearly defined as per Validation Master Plan

## Slide 17: TYPES OF VALIDATION
TYPES OF ANALYTICAL METHOD FOR VALIDATION:
Identification Tests: To identify 1 or more analyte in sample, Property of test samples compared with standard samples
Quantitative & limit test for impurities: To reflect the purity of a sample.
Quantitative Tests (assay): To measure analyte present in given sample

## Slide 18: TYPES OF VALIDATION
VALIDATION CHARACTERISTICS OF ANALYTICAL METHOD VALIDATION:
Specificity
Linearity
Range
Accuracy
Precision: Repeatability, Intermediate precision, reproducibility
Detection Limit
Quantitation Limit
Robustness
System suitability test

## Slide 19: TYPES OF VALIDATION
1. Specificity: ability of method to assess the analyte in presence of other components (degradants, impurities, matrix etc.) of sample.
Specificity of method should be determined during validation of Identification test, impurity testing, assay methods.
Specificity is essential to ensure correct identity of a sample, that the method provides an accurate statement of content of impurities etc. & to provide exact estimate of the potency of the analyte in sample.
If impurity or degradation products are available, then assay methods should demonstrate differences in analyte in presence of impurities. Pure drug substance is spiked with required levels of impurities or excipients & compare results obtained on unspiked samples.
If not available, the specificity of method can be demonstrated by comparing test results of samples containing impurities/degradation products to pharmacopoeial or other validated analytical method.

## Slide 20: TYPES OF VALIDATION
2. Linearity of analytical method
It is the ability to produce test results within a given range directly proportional to the amt/conc. of analyte in the sample. Dilution of stock solution of drug can be used. Evaluation of linearity is done by visual inspection of plots which are function of amt/conc. of analyte
By using statistical methods, results can be evaluated.
Ex. Regression line can be calculated by least square method, and regression line can be used to calculate degree of linearity
Slope of regression line, y intercept, correlation coefficient & residual sum of squares along with the plot of data are estimated
Minimum 5 conc. are recommended for obtaining linearity

## Slide 21: TYPES OF VALIDATION
3. Range: It is the interval between the upper & lower conc. of analyte in sample for which suitable level of precision, accuracy & linearity has been done. It is mainly obtained from linearity studies
Following minimum specified ranges for consideration:
For assay of drug subs./finished product- 80%-120% of test conc.
For content uniformity : 70-130% of test conc. According to dosage forms ex: metered dose inhalers
For dissolution testing: ± 20% over the specified range
For impurities known to be extremely potent or to produce toxic or unexpected pharmacological effects: Detection/quantitation limit should be in proportion to level at which impurities can be controlled
If assay & purity are done together as one test & only a 100% std. is used, linearity should cover the range from reporting level of impurities to 120% of the assay spec.

## Slide 22: TYPES OF VALIDATION
4. Accuracy: “known as trueness” Agreement between true and measured value. Minimum 9 determinations of 3 conc. levels is done to check accuracy. It should cover the specified range.
ICH defines the accuracy of an analytical procedure as the closeness of agreement between the values that are accepted as conventional true values or an accepted reference value and the value found.
Accuracy is usually reported as % recovery by the assay (using the proposed analytical procedure) of known added amount of analyte in the sample or as the difference between the mean and the accepted true value together with the confidence intervals.
Typical accuracy of the recovery of the drug substance is expected to be about 99–101%.
Typical accuracy of the recovery of the drug product is expected to be about 98–102%.

## Slide 23: TYPES OF VALIDATION
4. Precision: Repeatability, Intermediate precision, reproducibility
The precision of an analytical procedure expresses closeness of agreement between a series of measurements
Precision may be considered at three levels: repeatability, intermediate precision and reproducibility
The precision of an analytical procedure is usually expressed as the variance, standard deviation or coefficient of variation of a series of measurements.
Repeatability: measure of the precision under the same operating conditions over a short interval of time, that is, under normal operating conditions of the analytical method with the same equipment. It is intra assay precision.
Intermediate Precision: variation within the same laboratory. Typical parameters that are investigated include day-to-day variation, analyst variation, and equipment variation. Experimental design will minimize the number of experiments that need to be performed
Reproducibility: means closeness of agreement in test results obtained b/w different laboratories using same analytical method, assessed based on inter laboratory trial. Methods in pharmacopoeias include reproduce.

## Slide 24: TYPES OF VALIDATION
5. Detection Limit/ limit of detection (LOD)
It is capability of an analytical method to detect the lowest amt of analyte in a sample. But quantitation of that amt is not considered.
Various approaches are there based on instrumental or non instrumental procedures. They are
a) Visual evaluation : minimum level at which analyte can be detected reliably is established.
b) Signal to noise approach: suitable with procedures with base line noise. Signal to noise ratio: measures signals from sample with known low conc. of analyte & compare the measured signals with blank samples, & establish minimum conc. At which analyte can be detected. Accepted limit is 3:1 or 2:1
c) Standard deviation of response & slope: Detection limit by this is given by DL= 3.3 SD/S, SD is std. deviation, S is slope of calibration curve. SD can be determined from SD of the blank or from calibration curve.

## Slide 25: TYPES OF VALIDATION
6. Quantitation Limit (QL/ LOQ limit of quantification) of an analytical procedure
It is lowest amount of analyte in a sample which can be quantitatively determined with suitable precision and accuracy. Used for the determination of impurities and/or degradation products.
Approaches based on instrumental or non instrumental procedures are
a) Visual evaluation : Minimum level at which analyte can be quantified reliably is established by taking sample with known conc.
b) Signal to noise approach: suitable with procedures with base line noise. Signal to noise ratio: measures signals from sample with known low conc. of analyte & compare the measured signals with blank samples, & establish minimum conc. At which analyte can be quantified. For QL signal is to noise ratio accepted limit is 10:1
c) Standard deviation of response & slope: QL is given by QL= 10 SD/S, SD can be determined from SD of the blank or from calibration curve.

## Slide 26: TYPES OF VALIDATION
7. Robustness: Robustness of an analytical procedure is a measure of the analytical method to remain unaffected by small but deliberate variations in method parameters and provides an indication of its reliability during normal usage. The evaluation of robustness is normally considered during the development phase and depends on the type of procedure under study. Experimental design (e.g., fractional factorial design or Plackett–Burman design) is common and useful to investigate multiple parameters simultaneously. Common method parameters that can affect the analytical procedure should be considered based on the analytical technique and properties of the samples:
1. Sample preparation: Extraction time , Sample solvent, Membrane filters, Sample and standard stability
2. HPLC conditions: Mobile-phase composition, Column used (equivalent columns, lots and/or suppliers, age), Temperature, Flow rate
3. Gas chromatography (GC) conditions: Column used, Temp. Flow rate

## Slide 27: TYPES OF VALIDATION
8. System suitability test: It is an integral part of many analytical procedures. The test are based on concept that the samples to be analyzed, analytical operations, equipment & electronics constitute an integral system that can be evaluated as such. It ensures that system is working properly at the time of analysis. Determination mode are repeatability, tailing factor T, capacity factor K’, resolution R, theoretical plates N

## Slide 28: VALIDATION MASTER PLAN
It is key document to implement validation programme in GMP regulated Pharm. Industry.
It summarizes the company’s overall philosophy, intentions, approaches to be used.
It is approved written plan & summary doc. Provides overview of entire validation operation.
It outlines the principles involved in qualification of a facility, defines the areas & systems to be validated & provides a written program for achieving & maintaining a qualified facility with validated processes.
Format & content of validation master plan includes:
Introduction: Validation policy, scope, location, schedule
Organizational structure: personnel responsibilities

## Slide 29: VALIDATION MASTER PLAN
3. Plant/process/product description: rational for inclusions or exclusions & extent of validation specific processes considerations that are critical & those requiring extra attention.
4. List of products/processes/systems to be validated, summarized in a matrix format, validation approach.
5. revalidation activities, actual status and future planning
6. Key acceptance criteria
7. Documentation format
8. Reference to the existing documents like required SOP’s
9. Time plans of each validation project and sub project

## Slide 30: CALIBRATION
It is process of comparing a reading on one piece of equipment or system with another piece of equipment that has been calibrated & referenced to a known set of parameters.
As per WHO: Calibration is set of operations that establish, under specified conditions, the relationship between values indicated by an instrument or system for measuring weight, temp. pH etc. , recording, & controlling or the values represented by a material measure, & the corresponding known values of a reference standard.
As per ICH: It is demonstration that a particular instrument or device produces results within specified limits by comparison with those produced by a reference or traceable standard over an appropriate range of measurements.

## Slide 31: Differences Between Calibration & Validation
CALIBRATION | VALIDATION
Purpose: to demonstrate that a particular instrument or device produces results with in specified limits by comparisons with those produced by reference over appropriate range of measurements | Purpose: to documented verification that provides high degree of assurance that a specific process, equipment, method or a system consistently produces a result meeting predetermined acceptance criteria
Performance of instrument or device is compared against a reference standard | No such reference std. are used
Ensures that instrument/ measuring device produce accurate results | Provides documented evidence that a process, equipment, method or system produces consistent results i.e. uniform batch are produced
Performed periodically, to identify drift of the measuring device or equipment & make them accurate | No such req. performed when changes or modifications happen to the existing system or once revalidation period is reached
Performed as per calibration SOP | Performed as per validation protocol
Adjust the precision & accuracy of instrument | Confirms precision & accuracy of equipment
Helps to eliminate bias | It confirms that no bias exist in instrument

## Slide 32: Qualification of UV visible spectroscopy
1. DESIGN QUALIFICATION
Documented evidence that the plant design agrees with design specs of costumer
3. OPERATIONAL QUALIFICATION
Documented evidence which shows that all parts of plant & equipment work within their specifications & process parameters
2. INSTALLATION QUALIFICATION
Written evidence that all parts of equipment are installed according to equipment and purchase spec.
4. PERFORMANCE QUALIFICATION
Documented evidence that all parts of a plant & processes produce products of specified quality under conditions of normal production for a longer period of time
Qualification is an act or process to assure complies with standard

## Slide 33: Qualification of UV visible spectroscopy
Installation procedure: While the UV instrument was shipped after precise adjustment & inspection at factory, it is recommended to install according to following procedure so as to provide its optimum performance & to meet user’s demand:
Installation site:
Room Temp. during use : 15-35C
Out of direct sunlight
No strong vibration or continuous weak vibration
No strong magnetic fields or electromagnetic fields
Humidity of 45 to 80%
No corrosive gases or organic or inorganic gases with absorptivity in UV range
Small amount of dust

## Slide 34: Qualification of UV visible spectroscopy
ACCEPTANCE PROCEDURE
Items to be checked | Specifications
Appearance | No defect
Number of parts | No missing parts
ROM check | Latest version
Linearity of Absorbance | Bent : ±0.002Abs Shock noise: ±0.004 Abs
Noise level | Noise width: ±0.002Abs Shock noise: ±0.004 Abs
Accuracy of wavelength | ±0.5nm
Repeatability of wavelength | ±0.1nm

## Slide 35: Qualification of UV visible spectroscopy
OPERATIONAL AND PERFORMANCE QUALIFICATION
Wavelength accuracy: it is defined as the deviation wavelength reading at an absorption band and emission band from the wavelength of the band
Acceptance: ±nm in UV range (200-380nm) &
± nm in visible range (380-800nm)
Three repeated scan of same peak should be within ±0.5nm
Stray Light: It is defined as the detected light of any wavelength that is outside the band width of the wavelength selected
Acceptance: the transmittance of the solution in a 1cm cell should be less than 0.01 or the absorbance value should be greater than 2
Resolution Power: the resolution power of UV-Vis Spectrophotometer is related to its spectral band width. The smaller the band width the finer the resolution. Spectral band width depends on the slit width & dispersive power of the monochromator

## Slide 36: Qualification of UV visible spectroscopy
Acceptance : The ratio of absorbance at 269nm & absorbance at 266 nm should be greater than 1.5
Noise: Noise is measurement affecting the accuracy at both end of absorbance scale. Photon noise from the light source affects the accuracy of the measurement leads to low absorbance
Acceptance : The RMS noise should be less than 0.001AU
Baseline flatness: the flat baseline test demonstrates that the ability of the instrument to normalize the light intensity measurement & spectral output at different wavelength throughout the spectral range
Acceptance: The measurement is typically less than 0.01AU
Stability: The lamp intensity is function of lamp age, temp. fluctuation, & wavelength of the measurement. These changes can lead to errors in the value of measurements, over an extended period of time
Acceptance : The deflection is less than 0.002AU/hr

## Slide 37: Qualification of UV visible spectroscopy
Photometric accuracy: it is determined by comparing the difference between the measured absorbance of the reference material & the established value.
Acceptance: six replicate measurements of the 0.006% w/v of potassium dichromate solution at 235, 257, 313 & 350nm should be less than 0.5% RSD.
Linearity: The linearity dynamic range of measurement is limited by stray light at high absorbance and by noise at low absorbance.
The accuracy of the quantification of the sample depends on the precision & linearity of the measurements.
Acceptance: Correlation coefficient: r ≥0.999